# Engineering change record

Cumulative changes from 2026-09-09 through 2026-09-14

Revision date: 2026-09-14

Scope This is the clean non-code record of the September 9–14 implementation line. It includes the September 9 metadata/layout/refactor work, September 10 timing and Andor streaming changes, September 11 lab-driven stream/settings fixes, and the September 14 timing-retention and capability corrections.

# Change chronology

| Date | Target | Implementation | Evidence/status |
| --- | --- | --- | --- |
| 2026-09-09 | Magnetic-field applicability | Record field now captures applicability at acquisition. None means not applicable; 0 mT remains a valid measured/entered value. Filename, ordinary CSV, gated series, and monitor use the captured record rather than the later widget state. | Source/offline complete; physical 0/nonzero/off matrix remains. |
| 2026-09-09 | Gated planning | Replaced materialized action lists with a constant-size lazy plan, exact count/memory guards, compact semantic preview, and preflight reuse. | 1,001-frame example planned in ~0.041 ms in review runtime. |
| 2026-09-09 | Gated averaging | Added incremental per-trace mean, sample SD, timing distributions, partial accepted output, one-file series saving, and pump/frame/repeat metadata. | Offline round trips; lab multi-frame gated runs passed. |
| 2026-09-09 | Acquisition controls | Compacted common controls, added progress reporting for long exposure/averaging, selected Newport channels, wrong-channel warning, and multiple smoothing kernels/boundaries/previews. | Offline complete; visual/physical checks remain. |
| 2026-09-09 | Metadata | Introduced conditional spectrum schema v4: grouped fields, canonical raw counts, processed column only when needed, requested/applied distinctions, strict background signature, Andor readback and runtime provenance. | Backward loading retained for v1–v4 and plain two-column files. |
| 2026-09-09 | Extensibility design | Defined capability registry, manifest/catalog, host resource leases, declarative procedures, supervised third-party execution, and a read-only field-sensor pilot roadmap. | Design only; no untrusted extension execution added. |
| 2026-09-10 | Andor exposure truth | Separated requested exposure, SDK-applied exposure, readout, keep-clean, and kinetic period. Applied exposure is primary saved integration; differing request is conditional. | Source complete; short mode-5 matrix remains. |
| 2026-09-10 | Andor throughput | Moved rapid owners to one armed SDK2 Run-Till-Abort stream with circular-buffer order, session keys, pre-request draining, explicit indices, and an unsaved gated preparation frame. | Lab showed ~60 Hz at 15 ms and ~19.95 Hz at 50 ms. |
| 2026-09-10 | Andor units/defaults | Uses the enumerated vertical speed nearest 16 µs/pixel by default; labels 0.1 MHz correctly as 100 kHz; exposes applied/readout timing separately. | Default/labels confirmed in lab. |
| 2026-09-10 | Background/I/O | Cached strict signatures and resampled backgrounds; automatic spectrum/series formatting and writes moved to one ordered executor. | Background subtraction throughput lab-confirmed near unsubtracted Live. |
| 2026-09-10 | Instrument behavior | Kept sources independently real/emulated/disconnected, GUI-first expected connection errors, automatic/explicit spectrometer choice, restartable Newport process, and QEPro post-change buffer synchronization. | Newport reconnect later passed; QEPro freshness still bench-gated. |
| 2026-09-11 | Andor stale events | Made GetNumberNewImages the readiness authority after a GUI freshness drain. Successful stale waits are ignored inside the original timeout; GetOldestImage runs only with unread data. | Gap probe 100/100; GUI Live and gated repeated acquisition passed. |
| 2026-09-11 | Gated midpoint selection | Passes an absolute exposure-midpoint target into the Andor stream and consumes early candidates separately from pre-request backlog. | Aligned 2026-09-14 probe confirms all 99 steady-state frames eligible. |
| 2026-09-11 | Spectrometer Details | Added explicit Apply and verify transaction through runtime/worker, stream stop, validation, ordered camera/Kymera changes, motion readback, calibration refresh, retryable failure, and a reversible controls probe. | Supported control matrix verified 2026-09-14. |
| 2026-09-11 | TEC session state | Caches a successful iDus target so the dialog does not visually reset while the connected camera remains at that target. | Source complete; cooler exit dialog defects remain. |
| 2026-09-14 | Fixed-window retention | Added Retain N inside fixed window for averaged delayed/interleaved runs: per-trace accepted target, start/midpoint reference, ± tolerance, finite attempt multiplier, immediate rejection release, partial failure output, and progress accounting. | 145/145 offline invocations; hardware validation required. |
| 2026-09-14 | Timing provenance | Added software transition basis and timing-reference fields. Averaged series schema v5 records tolerance, evaluated/retained/rejected counts, attempt ceiling, detector basis, transition basis, and signed-error definition. | Round-trip/serialization tests pass. |
| 2026-09-14 | Capability hardening | Production FVB schema now offers binning 1 only after binning 4 returned DRV_NOT_SUPPORTED. Focus remains visible but read-only after 423→400 returned ATSPECTROGRAPH_P2INVALID. Probe emits full control schema. | Evidence-based conservative boundary; diagnosis remains. |
| 2026-09-14 | Low-light roadmap | Specified optical-throughput optimization, detector characterization, paired phase cycling, variance propagation, robust cosmic-ray handling, wavelength-band/template estimators, and physical timing calibration. | Design documented; paired workflow not implemented. |

# Architecture changes

AndorSDK2Camera owns stream state, buffer freshness, frame dating, target eligibility, and driver indices; the composite adapter remains responsible for calibrated spectra and user-facing state.

The gated planner remains pure/lazy. Retention adds a bounded attempt dimension without materializing actions or retaining rejected spectra.

RetainedTraceGoal counts accepted frames by stable trace identity; GatedSeriesAccumulator incrementally stores accepted mean/variance and timing samples.

Individual spectrum schema remains v4. Averaged gated series advances to v5 because retention changes the run-level timing/attempt contract.

Capability exposure is evidence-gated: a queryable value is not automatically assumed writable.

# Validation performed

Uploaded 2026-09-14 ZIP passed archive integrity/path checks; 13 JSON and eight CSV files were fully parsed.

All 141 Python files compile under the available CPython 3.12 runtime.

145 of 145 hardware-independent compatibility invocations pass.

No duplicate top-level definitions, trailing whitespace, or Python lines over 100 characters were found.

A fake-PySide coordinator integration retained exactly two valid spectra from four attempts under a six-attempt ceiling.

PySide6, Python 3.14, and vendor DLL execution were unavailable in the review environment; listed hardware checks are not replaced by these results.

# Files materially changed in the 2026-09-14 pass

| File(s) | Reason |
| --- | --- |
| core/gated_acquisition.py | Retention settings, validation, timing reference and transition basis metadata. |
| processing/gated_timing.py | Fixed two-sided guard and per-trace retained-target tracker. |
| planning/gated_sequence.py | Target-versus-attempt accounting, lazy attempt ceiling, preflight warnings. |
| controllers/gated_acquisition_coordinator.py | Timing selection, accept/discard/retry state, partial/final output metadata. |
| panels/gated_acquisition_panel.py; panels/main_window.py | Controls, persisted preferences, forced compatible options, preview/status. |
| processing/gated_averaging.py; io_utils/gated_series_io.py | Accepted-only aggregation and schema-v5 timing/attempt groups. |
| io_utils/spectrum_io.py | Transition-basis/timing-reference spectrum metadata round trip. |
| devices/andor_sdk2.py; devices/andor_spectrograph.py; devices/andor_adapter.py | Conservative FVB binning/focus capabilities and revision provenance. |
| dialogs/spectrometer_details_dialog.py; troubleshooting/andor_controls_probe.py | Disable unsupported editors; emit full capability schema. |
| README.md; ARCHITECTURE.md; REVIEW_20260909.md; ANDOR_TIMING_ANALYSIS_20260910.md; memory/docs | Cumulative product, architecture, evidence, and task records. |