# Plugins and user-defined procedures

Feasibility, boundaries, alternatives, security model, and staged roadmap

| STATUS Reviewed source design | REVISION 2026-09-10 | AUDIENCE Lab users and maintainers |
| --- | --- | --- |

| RECOMMENDATION | Both features are feasible. Use versioned capability contracts, trusted built-ins, supervised subprocesses for third-party Python, and a declarative procedure language that can act only through host arbitration/runtime services. |
| --- | --- |

# Executive decision

The application already has useful extension seams—device adapters, workflow coordinators, a sequence arbiter, worker threads, and versioned records. It is not yet safe to expose those classes directly to arbitrary code. The next step is a small service registry and stable contracts, not an integrated IDE or an unrestricted plugins folder.

An instrument plugin implements one or more known capabilities; it does not orchestrate unrelated devices or call MainWindow.

A measurement plugin uses public services to add a coordinator, analysis, export, or optional trusted panel.

A procedure is a validated, serializable recipe of registered steps—not arbitrary Python in the GUI process.

Advanced Python automation can later run out of process through a narrow, permissioned API.

# Terms and boundaries

| Concept | Means here | Explicitly not |
| --- | --- | --- |
| Instrument plugin | Discovers/connects one family and provides typed spectrometer, meter, laser, field, camera, or trigger capabilities. | Direct MainWindow calls, cross-instrument orchestration, or unbounded import-time threads. |
| Measurement plugin | Registered workflow/analysis/export factory that consumes capabilities through host services. | Raw access to another plugin's device handle or replacement of the global arbiter. |
| UI contribution | A schema-driven panel/dialog or, for trusted code, a QWidget factory bound to a service. | Monkey-patching existing widgets or owning vendor threads from event handlers. |
| Procedure | Validated steps such as configure → laser ON → wait → acquire → repeat → save. | Python exec, hidden OS access, or a hard-real-time guarantee. |
| Procedure runner | State machine that resolves resources, preflights, executes/cancels, reports progress, and guarantees cleanup. | A general IDE or substitute for physical interlocks. |
| Plugin registry | Versioned manifests, factories, capability schemas, metadata namespaces, and optional UI declarations. | An untyped global dictionary or automatic device connection during discovery. |

# Plugin architecture options

| Option | Structure | Advantages | Disadvantages | Fit |
| --- | --- | --- | --- | --- |
| A. In-process entry points | Package metadata exposes factories loaded into the GUI process. | Lowest cost and latency; easy built-in migration; direct NumPy arrays. | Crash/deadlock/dependency or malicious code affects GUI; no safe unload. | Signed/trusted first-party only. |
| B. Manifest + subprocess RPC | Read manifest first; execute plugin in spawned worker; exchange typed messages/shared arrays. | Fault/dependency isolation, restart, timeouts, permissions, safer third parties. | Protocol/version work and shared-memory lifecycle; modest latency. | Recommended third-party default. |
| C. External service | Separate local/remote service over a versioned transport. | Other languages/OSes, remote hardware, strongest separation, reusable outside GUI. | Deployment, discovery, authentication, and network failure complexity. | Facility/remote option later. |
| D. Scan/import .py folder | Import source files found in a user directory. | Very fast prototype. | No dependency, version, safety, or security boundary; fragile diagnostics. | Do not ship. |

## Recommended hybrid

Make built-in adapters conform to the public contracts, but keep them in-process initially.

Discover manifests without importing code. Run explicitly trusted first-party packages in-process and third-party packages under a supervisor.

Use shared memory/memory maps for spectra and images; use small typed messages for commands and status.

Prefer schema-generated controls. Allow custom QWidget code only for trusted packages.

Keep leases, laser cleanup, global logging, and file provenance in the host.

# Capability contracts

| Capability | Core operations | Boundary / returned evidence |
| --- | --- | --- |
| Spectrometer | Identity/capabilities, configure, acquire, timing, health, close. | Arrays, applied settings, and timing basis; never power/laser calls. |
| Power meter | Identity/channels, read, status decode, health, close. | Timestamped values/status; host owns selected channel roles. |
| Laser source | Channels, setpoint, enable/disable, readback. | Host-enforced disable/cancel and explicit interlock semantics. |
| Field sensor/controller | Read field/uncertainty; optionally set target. | Distinguishes manual field from measured readback/provenance. |
| Imager/beam profiler | Configure, acquire image, calibration. | Large arrays via shared memory; analysis separate from acquisition. |
| Trigger/timing | Arm, fire, timestamps/clock metadata, abort. | Required before claiming hardware-timed procedures. |
| Analysis/export | Consume immutable records; emit derived result/artifact. | No direct device access; versioned deterministic parameters. |

# Manifest and lifecycle

A manifest must be readable before executable import. It should contain a stable ID, display name, semantic version, compatible host API range, entry point/command, capabilities, permissions, configuration schema, metadata namespace, and optional UI declarations.

{
  "id": "org.example.field_meter",
  "version": "1.2.0",
  "host_api": ">=1.0,<2.0",
  "runtime": "subprocess-python",
  "provides": ["field_sensor.v1"],
  "permissions": ["serial:COM12", "filesystem:none"],
  "config_schema": "schemas/config-v1.json",
  "metadata_namespace": "org.example.field_meter"
}

Discover and validate manifests; show incompatible/disabled plugins without executing them.

Resolve configuration and concrete permissions; require explicit enablement for third-party code.

Start worker, negotiate API/protocol versions, and request capability schema.

Connect only on session/user request; normalize connection states and structured errors.

Lease resources through the host; commands carry IDs, deadlines, cancellation, and idempotency expectations.

On crash/timeout, retire worker, release leases, invoke host fail-safe actions, and preserve diagnostics.

On shutdown, stop procedures, disable hazardous outputs, close adapters, then terminate workers.

# Procedure implementation options

| Option | Structure | Advantages | Disadvantages | Fit |
| --- | --- | --- | --- | --- |
| Declarative JSON/YAML DSL | Typed steps, variables, bounded loops/conditions, waits, acquire/save; validated before run. | Reviewable, serializable, reproducible, preflightable, safe provenance. | Needs a schema/editor and explicit step vocabulary. | Recommended v1. |
| Visual workflow builder | Graphical editor emits the same declarative DSL. | Accessible, discoverable, live validation. | Large UI project; complex graphs can become unwieldy. | After DSL stabilizes. |
| Restricted expressions | DSL permits bounded expressions for derived values/conditions. | More flexibility without full Python. | Evaluator security, units, and typing complexity. | Narrow v2 addition. |
| Sandboxed Python SDK | Separate user process calls a narrow host API. | Maximum scientific flexibility and normal tooling. | Desktop sandbox limits; harder cancellation, dependencies, review, and reproducibility. | Advanced opt-in. |
| In-process Python / IDE | Edit and execute Python inside GUI. | Fast demo; familiar to programmers. | Can bypass every safety/data boundary and freeze/corrupt GUI. | Do not implement first. |
| External automation client | Documented local API used from VS Code/Jupyter/other run engine. | No IDE maintenance; mature tooling; flexible language. | Two-process lifecycle, auth, and resource UX. | Strong IDE alternative. |

## Recommended procedure model

Compile a user definition into the same action/state concepts as built-in coordinators. Preflight resolves capabilities, validates units and ranges, estimates frames/time/memory/files, checks resource conflicts, and creates a compact preview. Execution is event-driven and never sleeps in the GUI thread.

procedure: delayed-spectrum-average
version: 1
requires: [spectrometer.v1, laser_source.v1]
parameters:
  repeats: {type: integer, minimum: 1, default: 100}
steps:
  - acquire_lease: [spectrometer, laser]
  - repeat: ${repeats}
    do:
      - laser.set: {enabled: true}
      - wait: {duration: 50 ms}
      - laser.set: {enabled: false, mark_transition: true}
      - spectrum.acquire_at: {delay: 0 ms, output: accumulator}
  - accumulator.save: {format: gated_series_csv}
finally:
  - laser.set: {enabled: false}
  - release_all: true

| TIMING | A procedure runner can minimize preparation and ordering overhead, but desktop software does not provide hard-real-time timing. Trigger-capable plugins must publish clock, edge, and uncertainty semantics explicitly. |
| --- | --- |

# Host structure and communication

| Component | Role | Key rule |
| --- | --- | --- |
| ExtensionCatalog | Reads manifests, compatibility, trust, and enable state. | No imports or connections during discovery. |
| PluginSupervisor | Process lifecycle, heartbeat, timeout, crash report, shared-memory cleanup. | One supervised session per plugin instance. |
| CapabilityRegistry | Maps stable capability IDs to typed proxies/factories. | No MainWindow references; negotiate versions. |
| ResourceManager | Named leases for spectrometer, laser channel, meter, camera, trigger. | Hazardous resources have host cleanup. |
| ProcedureCompiler | Schema/units/limits validation, step resolution, cost preview. | Pure preflight; no hardware side effects. |
| ProcedureRunner | Event state machine, progress, cancellation, retry/checkpoint/finally. | Calls capability proxies only. |
| MetadataRegistry | Namespaces fields and captures extension/procedure versions and hashes. | No collisions with core schema. |
| UIContributionHost | Builds schema controls and binds service commands. | Custom native widgets trusted-only. |

# Safety, trust, and failure rules

Plugins never receive raw objects from another plugin or a MainWindow reference.

Laser enable requires a named lease and host policy. Abort, disconnect, crash, procedure failure, and app shutdown invoke host-owned disable-all sequencing.

A timeout retires an unresponsive process; never attempt unsafe thread termination.

Permissions are deny-by-default and concrete: serial port, vendor runtime/USB, network endpoint, selected folder, or none.

Validate RPC on both sides: finite numbers, array shapes, units, command IDs, protocol version, and stale responses.

Record definition hash/version, resolved plugin versions, parameters, outcome, rejected frames, and cleanup result once metadata fields are approved.

Describe Python execution on Windows as process isolation plus permissions—not as a perfect security sandbox.

# Roadmap

| Stage | Work | Exit criterion |
| --- | --- | --- |
| 0 — Prerequisites | GUI integration tests; split MainWindow connection/acquisition/workflow presenters; structured error/cancel contracts. | Stable seams without behavior change. |
| 1 — Contracts | Protocol/ABC package for capabilities, identity, timing, configuration, health, progress, and close; migrate built-ins. | Built-in registry, no third-party loading. |
| 2 — Procedures | v1 JSON Schema, units, step registry, compiler, cost preview, runner, finally cleanup, progress, provenance; parity-convert one gated workflow. | Trusted built-in declarative procedures. |
| 3 — Plugin catalog | Manifest/API compatibility, namespaces, enable UX, diagnostics, install policy, conformance suite. | Discovery without import. |
| 4 — Subprocess RPC | Supervisor, framed messages, heartbeat/deadlines, shared arrays, restart, permissions, host cleanup. | Isolated third-party instrument pilot. |
| 5 — Authoring UX | Schema-aware editor/form, validation, compact preview, emulator dry run, templates, run history. | Usable authoring without full IDE. |
| 6 — Advanced | Trusted QWidget contributions, analysis/export plugins, external API, restricted expressions, optional Python SDK. | Broader ecosystem after safety matures. |
| 7 — Governance | Signing/trust labels, dependency pins, compatibility matrix, update/rollback, docs/examples. | Maintainable multi-user distribution. |

# Alternatives that may solve needs sooner

| Alternative | Value | Tradeoff |
| --- | --- | --- |
| Configuration-driven SCPI/serial adapters | Very quick for simple read/set instruments with no custom package. | Poor for binary SDKs, streaming, complex state, and safety logic. |
| Semantic macro recorder | Records host commands and replays them; discoverable and no programming. | Limited branching; brittle if it records widgets instead of commands. |
| External local API | Users automate from a normal IDE/Jupyter without embedding one. | Requires API lifecycle/version/auth and the same arbitration. |
| Bridge a lab automation framework | Mature plans, queues, data catalogs, and ecosystem. | Integration/migration cost; vendor and safety adapters still needed. |
| Measurement templates only | Lowest risk/support burden for common experiments. | Cannot cover novel devices/logic; not a community extension model. |

| NEXT DECISIONS | Trust, first pilot, local versus remote scope, procedure expressiveness, authoring UX, and provenance questions are collected in QUESTIONS_20260910.docx. |
| --- | --- |

# Minimum viable extension pilot

Define field_sensor.v1 and implement one read-only built-in adapter plus a fake conformance adapter.

Register built-ins through the new capability registry while preserving today's direct behavior as a regression baseline.

Compile one declarative procedure: read field → acquire spectrum → save; preflight units, availability, ownership, and output count.

Run the same procedure entirely against emulators, then with only the real field sensor, then with the real selected spectrometer.

Record structured progress, cancellation, failure, and final cleanup; do not include custom QWidget or arbitrary Python in the pilot.

## Pilot acceptance criteria

Discovery performs no plugin import or hardware I/O.

A missing/crashed field sensor cannot stop spectrum acquisition or freeze the GUI.

Procedure cancellation releases every lease and produces an explicit outcome.

The saved result distinguishes manual field, measured field, uncertainty, source identity, and read age after those metadata fields are approved.

The plugin/procedure API is versioned and covered by host-side contract tests before a second device type is attempted.

# Decisions incorporated on 2026 09 10

The extension runtime remains deferred, but the answered design constraints now narrow the first implementation. The host continues to own acquisition leases, hazardous output cleanup, metadata, cancellation, and audit logs.

| Area | Selected direction | Implementation consequence |
| --- | --- | --- |
| Install authority | Explicit administrative GUI mode such as --admin, combined with ordinary OS write permission. | Do not silently enable installation because a Windows account belongs to Administrators. Enabling an installed approved plugin can remain a normal user action. |
| First plugin | Read-only sensor, initially a magnetic-field instrument. | Small capability surface, no actuation, immediate metadata use, and a safe conformance pilot. |
| Authoring | Schema-aware text editor with keyword highlighting, generated parameter form, validation, and emulator dry-run preview. | One canonical procedure document feeds both text and form views; neither executes while being edited. |
| Remote scope | Network-ready contracts are likely required. | Use transport-neutral request/response/event models and avoid passing Python objects or local paths across the boundary. |
| Procedure language | Not yet selected. | Compare YAML and JSON source plus a typed internal plan before fixing v1 syntax. |
| External automation | Interaction model not yet selected. | Compare GUI-hosted localhost API, headless service, queued jobs, and notebook/IDE clients before exposing commands. |

# Procedure source and execution options

| Option | User interaction | Advantages | Disadvantages | Recommendation |
| --- | --- | --- | --- | --- |
| YAML source plus JSON Schema | Edit readable YAML; form edits the same model; validate/highlight continuously; dry-run compiles to a plan. | Readable units/comments; approachable; mature schema tooling. | Whitespace errors; YAML has surprising implicit types unless restricted. | Best human-authored v1 if parser uses a safe restricted subset. |
| JSON source plus JSON Schema | Edit strict JSON or use form; validate/highlight; dry-run compiles to the same plan. | Unambiguous types; simple interchange; strong tooling. | Verbose and less pleasant for long procedures; no comments in standard JSON. | Support as canonical interchange even if YAML is primary. |
| Dual YAML and JSON | Open either format; normalize to one typed internal representation; save in chosen format. | Meets human and machine needs without two runtimes. | More parser/round-trip tests and formatting decisions. | Recommended if maintenance cost is acceptable. |
| Visual graph builder | Drag registered steps, connect branches, edit forms, then inspect generated source/plan. | Accessible and discoverable. | Large UI project; graph/source synchronization; poor first prototype. | Add only after the step vocabulary stabilizes. |
| Python procedure SDK | Write code in VS Code/Jupyter and call a permissioned client API from a separate process. | Maximum flexibility and normal scientific tooling. | Harder reproducibility/cancellation; code can access the host OS; cannot be treated as safe declarative input. | Advanced opt-in after the external API is mature. |

All source formats should compile into one typed, immutable procedure plan. The compiler resolves units, capabilities, bounded loops, estimated time, frames, memory, files, and leases without hardware side effects. The runner then executes that plan as an event-driven state machine with progress, abort, timeouts, and a mandatory finally cleanup path.

# External API deployment and interaction options

| Model | How a user works | Advantages | Risks and cost | Best use |
| --- | --- | --- | --- | --- |
| GUI-hosted localhost API | Start GUI, enable API, connect from VS Code/Jupyter, inspect capabilities, submit a validated procedure, watch events, cancel. | Fastest path; reuses current runtime/arbiter and visible safety state. | Automation dies with GUI; port/auth/lifecycle UX; remote use requires care. | Local advanced-user pilot. |
| Headless measurement service plus GUI client | Start managed service at instrument PC; GUI/notebook/remote client authenticates and submits jobs. | Clean remote model; GUI restart does not kill service; multiple clients can queue. | Largest refactor; service deployment, ownership, authorization, upgrades, and physical-presence policy. | Recommended long-term remote architecture. |
| Queued job directory or broker | Export signed/validated procedure job; service accepts it and writes progress/results. | Simple disconnection tolerance and audit trail. | Not interactive; cancellation/status latency; file/broker security. | Batch facilities or intermittent links. |
| Direct plugin RPC only | Each remote instrument exposes capability RPC; local GUI still coordinates workflows. | Incremental remote devices and language independence. | Network jitter enters cross-device timing; distributed cleanup and clock semantics are difficult. | Remote non-time-critical sensors first. |
| Integrated Python IDE | Write and execute arbitrary Python inside the GUI. | Single-window convenience. | Can freeze UI, bypass safety/metadata, and creates a full editor/debugger/environment product. | Do not prioritize. |

Local API v1 should expose identity/capability discovery, procedure validation and dry run, job submission, progress/events, cancellation, result references, and connection state. It should not expose raw MainWindow or vendor objects.

Remote requests need authentication, TLS or an authenticated tunnel, named user/session identity, explicit resource leases, replay-resistant command IDs, deadlines, and a host-side laser policy.

Time-sensitive multi-instrument steps should execute on the instrument host. A remote client submits a plan; it must not issue millisecond laser and detector commands over the network.

Large spectra/images should use bounded streaming or shared/object storage references rather than embedding arrays in control messages.

# Revised implementation roadmap

| Stage | Implementation | Exit evidence |
| --- | --- | --- |
| One | Extract stable identity, connection, read-only sensor, spectrometer, power, laser, progress, error, and cancellation contracts from built-ins. | Built-ins pass contract tests without dynamic loading. |
| Two | Add ExtensionCatalog manifest validation and --admin install/remove flow; normal users enable only installed approved entries. | Discovery performs no import or hardware I/O; incompatible entries remain diagnosable. |
| Three | Implement field_sensor.v1 in-process for a trusted built-in and through a supervised fake subprocess. | Sensor loss cannot freeze acquisition; value/source/age are available to a record builder but not yet persisted without approval. |
| Four | Define restricted YAML/JSON Schema, typed internal plan, step registry, compiler, form/text synchronization, highlighting, and emulator dry run. | A read-field then acquire then save procedure validates and previews without side effects. |
| Five | Add ProcedureRunner using ResourceManager leases, current SequenceArbiter, progress, cancellation, retry policy, and finally cleanup. | Emulated and real read-only pilot runs are reproducible and cancellation releases all resources. |
| Six | Expose the same services through a permissioned localhost API and normal VS Code/Jupyter client library. | API cannot bypass leases, validation, metadata, or laser cleanup. |
| Seven | Generalize transport for remote read-only instruments, then evaluate a headless measurement service for remote actuation. | Authenticated failure/reconnect/upgrade/clock behavior passes an integration matrix. |
| Eight | Add third-party subprocess packages, signing/trust labels, dependency isolation, update/rollback, conformance tools, and documentation. | A plugin crash is contained and recoverable with complete diagnostics. |

The remaining decisions are collected in QUESTIONS_20260910.docx. No plugin loader, procedure executor, remote listener, or arbitrary Python execution is included in the current GUI source.