# Spectrum metadata catalog

Current schema, consolidation decisions, compatibility map, and proposed provenance

| STATUS Reviewed source design | REVISION 2026-09-11 / schema v4 | AUDIENCE Lab users and maintainers |
| --- | --- | --- |

| OUTCOME | Schema v4 retains the compact 09/09 layout and adds approved software provenance, readback-confirmed Andor state, continuous-buffer timing, selected Newport channels, background signature, smoothing boundary, equivalent provenance in one-file gated averages, and a distinct count of stream frames rejected to meet a gated exposure-midpoint target. |
| --- | --- |

# How to read this catalog

Names use group.field notation. On disk, a group occupies one CSV comment row containing key=value entries.

“Conditional” means the row or named value is omitted when not applicable. Omission is deliberate—not a false value.

Usefulness is classified as Essential, High, Analysis, Diagnostic, or Low/legacy.

Schema-v1 and schema-v2 files still load. Saving a loaded record emits schema v4.

# Schema-v4 individual-spectrum fields

file_type
Type / values: str: spectrum
Written when: Always
Plain-English meaning: Identifies the CSV payload family.
Use / value: Essential: selects parser.
Decision: Retained.

schema_version
Type / values: int: 4
Written when: Always
Plain-English meaning: Version of the current compact metadata layout.
Use / value: Essential: compatibility.
Decision: Retained.

timestamp_utc
Type / values: ISO-8601 str
Written when: Always
Plain-English meaning: Absolute UTC dating of the completed record.
Use / value: Essential: chronology/provenance.
Decision: Retained; replaces nonportable absolute perf-counter values.

spectrometer.backend
Type / values: str: qepro, andor, emulator, …
Written when: When identity exists
Plain-English meaning: Adapter family used.
Use / value: High: interpretation/reproduction.
Decision: Grouped.

spectrometer.model
Type / values: str
Written when: When identity exists
Plain-English meaning: Instrument/detector model description.
Use / value: High: provenance.
Decision: Grouped.

spectrometer.serial
Type / values: str
Written when: When available
Plain-English meaning: Detector/spectrometer serial number.
Use / value: High: provenance.
Decision: Grouped.

spectrometer.spectrograph_serial
Type / values: str
Written when: Andor composite only
Plain-English meaning: Kymera serial distinct from camera serial.
Use / value: High: composite identity.
Decision: Grouped.

spectrometer.max_intensity
Type / values: finite float, counts
Written when: When reported
Plain-English meaning: Digital/full-scale ceiling used during acquisition.
Use / value: Analysis: saturation context.
Decision: Grouped; retained instead of signal max.

acquisition.integration_ms
Type / values: finite float
Written when: Always
Plain-English meaning: Exposure actually applied by adapter/SDK.
Use / value: Essential: normalization and reproduction.
Decision: Primary value is applied, not merely requested.

acquisition.integration_requested_ms
Type / values: finite float
Written when: Only when it differs from applied
Plain-English meaning: User-requested exposure before device rounding/clamping.
Use / value: Diagnostic: explains device adjustment.
Decision: Collapsed into applied-only form when equal.

acquisition.call_duration_ms
Type / values: finite float
Written when: When controller bounds exist
Plain-English meaning: Wall time spent in adapter acquisition call.
Use / value: Diagnostic: throughput/latency.
Decision: Replaces separate monotonic start/finish rows.

acquisition.averages
Type / values: int ≥ 1
Written when: Always
Plain-English meaning: Number of detector frames requested for this returned spectrum.
Use / value: Essential: noise/statistical context.
Decision: Grouped.

acquisition.averaging
Type / values: str: software or device
Written when: When requested and applied modes match
Plain-English meaning: Actual averaging implementation.
Use / value: High: behavior/reproduction.
Decision: Single field replaces flag + option.

acquisition.averaging_requested
Type / values: str
Written when: Only on fallback/mismatch
Plain-English meaning: User-selected averaging implementation.
Use / value: Diagnostic.
Decision: Written with averaging_applied only when different.

acquisition.averaging_applied
Type / values: str
Written when: Only on fallback/mismatch
Plain-English meaning: Implementation actually used.
Use / value: High when mismatch occurs.
Decision: Replaces device_averaging_used flag.

detector_timing_ms.basis
Type / values: str enum-like description
Written when: When adapter supplies timing
Plain-English meaning: States what the timing bounds/estimate represent.
Use / value: Essential for gated interpretation.
Decision: Separate backend-timing group.

detector_timing_ms.exposure
Type / values: finite float, ms
Written when: When available from adapter
Plain-English meaning: SDK-applied exposure duration.
Use / value: High for timing/throughput.
Decision: Grouped; omitted if unavailable.

detector_timing_ms.readout
Type / values: finite float, ms
Written when: When available from adapter
Plain-English meaning: SDK-reported detector readout.
Use / value: High for timing/throughput.
Decision: Grouped; omitted if unavailable.

detector_timing_ms.keep_clean
Type / values: finite float, ms
Written when: When available from adapter
Plain-English meaning: SDK-reported keep-clean overhead.
Use / value: High for timing/throughput.
Decision: Grouped; omitted if unavailable.

detector_timing_ms.kinetic_cycle
Type / values: finite float, ms
Written when: When available from adapter
Plain-English meaning: SDK-reported minimum frame/kinetic period.
Use / value: High for timing/throughput.
Decision: Grouped; omitted if unavailable.

detector_corrections.electric_dark
Type / values: bool
Written when: Always
Plain-English meaning: Whether electric-dark correction was requested/applied by supported adapter.
Use / value: High: intensity provenance.
Decision: Combined with nonlinearity on one row.

detector_corrections.nonlinearity
Type / values: bool
Written when: Always
Plain-English meaning: Whether nonlinearity correction was requested/applied.
Use / value: High: intensity provenance.
Decision: Combined with electric dark on one row.

field_value_mT
Type / values: finite float; zero valid
Written when: Only when Record field checked
Plain-English meaning: Magnetic field deliberately associated with spectrum.
Use / value: High when field-dependent.
Decision: None/unchecked omits it everywhere; 0 is retained.

power.before_W
Type / values: JSON array of floats
Written when: When spectrum-associated power measured
Plain-English meaning: Per-channel power immediately before acquisition.
Use / value: High: bracketing/drift.
Decision: Combined with after_W; not reduced to a mean.

power.after_W
Type / values: JSON array of floats
Written when: When spectrum-associated power measured
Plain-English meaning: Per-channel power immediately after acquisition.
Use / value: High: bracketing/drift.
Decision: Combined with before_W.

power_status_abnormal.before_words
Type / values: JSON int array
Written when: Only if any status/power is abnormal
Plain-English meaning: Newport channel status words before.
Use / value: Diagnostic/safety.
Decision: Normal status omitted; abnormal evidence retained.

power_status_abnormal.after_words
Type / values: JSON int array
Written when: Only if any status/power is abnormal
Plain-English meaning: Newport channel status words after.
Use / value: Diagnostic/safety.
Decision: Normal status omitted; abnormal evidence retained.

power_status_abnormal.before_command
Type / values: int
Written when: Only if any status/power is abnormal
Plain-English meaning: Newport command status before.
Use / value: Diagnostic/safety.
Decision: Normal status omitted; abnormal evidence retained.

power_status_abnormal.after_command
Type / values: int
Written when: Only if any status/power is abnormal
Plain-English meaning: Newport command status after.
Use / value: Diagnostic/safety.
Decision: Normal status omitted; abnormal evidence retained.

run_identifier
Type / values: str
Written when: When user supplies one
Plain-English meaning: User grouping/name for the experimental run.
Use / value: High: organization.
Decision: Retained literal text.

notes
Type / values: str, including line breaks
Written when: When nonempty
Plain-English meaning: Free-form experimental note.
Use / value: High: human context.
Decision: One literal CSV field; notes_json removed.

scan.index
Type / values: int ≥ 0
Written when: Only during calibration/power scan
Plain-English meaning: Zero-based scan point.
Use / value: Analysis
Decision: Group presence replaces scan_active.

scan.count
Type / values: int ≥ 1
Written when: Only during calibration/power scan
Plain-English meaning: Total scan points.
Use / value: Analysis
Decision: Group presence replaces scan_active.

scan.basis
Type / values: str
Written when: Only during calibration/power scan
Plain-English meaning: Requested vs expected-power basis.
Use / value: High
Decision: Group presence replaces scan_active.

scan.spacing
Type / values: str
Written when: Only during calibration/power scan
Plain-English meaning: Linear/log/custom scan spacing.
Use / value: High
Decision: Group presence replaces scan_active.

laser.box_id
Type / values: str
Written when: Scan/gated frame or selected laser context
Plain-English meaning: Laser box identity.
Use / value: High for illumination provenance.
Decision: Grouped; transient COM port removed.

laser.channel
Type / values: int
Written when: Scan/gated frame or selected laser context
Plain-English meaning: Selected box channel.
Use / value: High for illumination provenance.
Decision: Grouped; transient COM port removed.

laser.wavelength_nm
Type / values: finite float
Written when: Scan/gated frame or selected laser context
Plain-English meaning: Nominal laser wavelength.
Use / value: High for illumination provenance.
Decision: Grouped; transient COM port removed.

laser.setpoint_W
Type / values: finite float
Written when: Scan/gated frame or selected laser context
Plain-English meaning: Commanded laser setpoint.
Use / value: High for illumination provenance.
Decision: Grouped; transient COM port removed.

power_scan.requested_W
Type / values: finite float, W
Written when: Only during power scan
Plain-English meaning: Requested sample power.
Use / value: High: calibration audit/reconstruction.
Decision: Retained conditionally despite ordinary-spectrum clutter concerns.

power_scan.expected_W
Type / values: finite float, W
Written when: Only during power scan
Plain-English meaning: Calibration-predicted actual power.
Use / value: High: calibration audit/reconstruction.
Decision: Retained conditionally despite ordinary-spectrum clutter concerns.

power_scan.filter_state
Type / values: str
Written when: Only during power scan
Plain-English meaning: Manual ND-filter state.
Use / value: High: calibration audit/reconstruction.
Decision: Retained conditionally despite ordinary-spectrum clutter concerns.

background.source_timestamp_utc
Type / values: ISO-8601 str
Written when: Only when background subtracted
Plain-English meaning: Identifies the compatible background used.
Use / value: High: subtraction provenance.
Decision: Grouped with the strict compatibility signature; group presence replaces background_subtracted.

smoothing.method
Type / values: str: none, boxcar, hann, hamming, gaussian, triangle, blackman, cosine, sine
Written when: Only when smoothing is active
Plain-English meaning: GUI-side smoothing kernel used.
Use / value: High for processed data.
Decision: Group presence indicates enabled.

smoothing.width
Type / values: int > 1
Written when: Only when smoothing active
Plain-English meaning: Kernel width in pixels.
Use / value: High for reproducibility.
Decision: Combined with method.

gated_frame.sequence_id
Type / values: UUID-like str
Written when: Only for individual gated frames
Plain-English meaning: Identifier joining frames in one sequence.
Use / value: Essential
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.mode
Type / values: str enum
Written when: Only for individual gated frames
Plain-English meaning: Gated plan type.
Use / value: Essential
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.index
Type / values: int
Written when: Only for individual gated frames
Plain-English meaning: Global frame index.
Use / value: Essential
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.count
Type / values: int
Written when: Only for individual gated frames
Plain-English meaning: Total planned frames.
Use / value: High
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.cycle
Type / values: int
Written when: Only for individual gated frames
Plain-English meaning: Pump/repeat cycle index.
Use / value: High
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.label
Type / values: str
Written when: Only for individual gated frames
Plain-English meaning: Human/column label.
Use / value: High
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.laser_state
Type / values: on/off str
Written when: Only for individual gated frames
Plain-English meaning: Requested state during frame.
Use / value: High
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.requested_delay_ms
Type / values: int
Written when: Only for individual gated frames
Plain-English meaning: Target delay after transition.
Use / value: Essential
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.phase
Type / values: int; -1 if unused
Written when: Only for individual gated frames
Plain-English meaning: Interleaved-decay phase.
Use / value: High
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.repeat
Type / values: int
Written when: Only for individual gated frames
Plain-English meaning: Zero-based repeat index.
Use / value: High
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_frame.repeats
Type / values: int ≥ 1
Written when: Only for individual gated frames
Plain-English meaning: Total repeat count.
Use / value: High
Decision: Grouped; sequence/phase retained because they reconstruct interleaving and joins.

gated_plan_ms.pump
Type / values: int ≥ 0, ms
Written when: Gated frames when applicable
Plain-English meaning: Requested laser-ON pump duration.
Use / value: High: requested timing provenance.
Decision: New in this revision.

gated_plan_ms.frame_gap
Type / values: int ≥ 0, ms
Written when: Gated frames when applicable
Plain-English meaning: Requested minimum gap between adjacent spectra.
Use / value: High: requested timing provenance.
Decision: New in this revision.

gated_plan_ms.repeat_gap
Type / values: int ≥ 0, ms
Written when: Gated frames when applicable
Plain-English meaning: Requested minimum delay between repeats.
Use / value: High: requested timing provenance.
Decision: New in this revision.

gated_call_timing_ms.request
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Coordinator request emission after transition.
Use / value: Diagnostic; start/request are key latency measures.
Decision: Kept separate from exposure timing to avoid implying call = exposure.

gated_call_timing_ms.start
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Adapter call start after transition.
Use / value: Diagnostic; start/request are key latency measures.
Decision: Kept separate from exposure timing to avoid implying call = exposure.

gated_call_timing_ms.midpoint
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Midpoint of adapter call bounds.
Use / value: Diagnostic; start/request are key latency measures.
Decision: Kept separate from exposure timing to avoid implying call = exposure.

gated_call_timing_ms.end
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Adapter call return after transition.
Use / value: Diagnostic; start/request are key latency measures.
Decision: Kept separate from exposure timing to avoid implying call = exposure.

gated_exposure_timing_ms.start
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Best estimated/bounded exposure-window start.
Use / value: Essential for time-resolved analysis.
Decision: Grouped.

gated_exposure_timing_ms.midpoint
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Best exposure midpoint estimate.
Use / value: Essential for time-resolved analysis.
Decision: Grouped.

gated_exposure_timing_ms.end
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Best estimated/bounded exposure-window end.
Use / value: Essential for time-resolved analysis.
Decision: Grouped.

gated_exposure_timing_ms.uncertainty
Type / values: finite float, ms
Written when: Gated frame when timing available
Plain-English meaning: Half-width/error bound of estimate.
Use / value: Essential for time-resolved analysis.
Decision: Grouped.

gated_exposure_timing_ms.basis
Type / values: str
Written when: Gated frame when timing available
Plain-English meaning: Method and semantics of estimate.
Use / value: Essential for time-resolved analysis.
Decision: Grouped.

gated_sample_windows_ms
Type / values: JSON array of [start,end] pairs
Written when: Gated frame with multiple averaged samples
Plain-English meaning: Per-average exposure windows relative to transition.
Use / value: High: exposes timing spread hidden by one mean.
Decision: Retained for software averaging.

gated_quality.status
Type / values: str
Written when: Gated frame when guard evaluates/provides value
Plain-English meaning: not_evaluated/warmup/accepted/outlier_flagged/outlier_discarded/timing_unavailable.
Use / value: High for inclusion decisions.
Decision: Grouped; status carries enabled/outcome semantics.

gated_quality.error
Type / values: finite float, ms
Written when: Gated frame when guard evaluates/provides value
Plain-English meaning: Residual from requested delay.
Use / value: High for inclusion decisions.
Decision: Grouped; status carries enabled/outcome semantics.

gated_quality.center
Type / values: finite float, ms
Written when: Gated frame when guard evaluates/provides value
Plain-English meaning: Rolling robust residual center.
Use / value: High for inclusion decisions.
Decision: Grouped; status carries enabled/outcome semantics.

gated_quality.robust_sigma
Type / values: finite float, ms
Written when: Gated frame when guard evaluates/provides value
Plain-English meaning: Rolling robust spread.
Use / value: High for inclusion decisions.
Decision: Grouped; status carries enabled/outcome semantics.

gated_quality.threshold
Type / values: finite float, ms
Written when: Gated frame when guard evaluates/provides value
Plain-English meaning: Applied rejection threshold.
Use / value: High for inclusion decisions.
Decision: Grouped; status carries enabled/outcome semantics.

snr.status
Type / values: str: ok or invalid
Written when: When SNR evaluation attempted
Plain-English meaning: Whether the stored estimate is usable.
Use / value: High for auto-tuning.
Decision: Replaces snr_valid boolean.

snr.reason
Type / values: str
Written when: Only when SNR invalid
Plain-English meaning: Human-readable invalidity reason.
Use / value: Diagnostic.
Decision: Combined with status; omitted on success.

snr_metrics.peak
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Peak SNR.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.integrated
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Integrated-band SNR.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.noise_sigma_counts
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Estimated noise sigma.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.peak_fraction
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Peak fraction of full scale.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.peak_signal_counts
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Peak baseline-subtracted signal.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.integrated_signal_counts_nm
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Integrated signal.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.integrated_noise_counts_nm
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Integrated noise.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.mean_signal_counts
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Mean signal-region counts.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.baseline_counts
Type / values: finite float
Written when: Only for valid SNR
Plain-English meaning: Baseline at signal center.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.signal_pixels
Type / values: int
Written when: Only for valid SNR
Plain-English meaning: Signal-region pixel count.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

snr_metrics.noise_pixels
Type / values: int
Written when: Only for valid SNR
Plain-English meaning: Noise-region pixel count.
Use / value: Analysis; peak/integrated are primary, rest diagnostic.
Decision: Grouped on one optional row.

wavelength_nm
Type / values: float column
Written when: Every pixel row
Plain-English meaning: Calibrated wavelength axis.
Use / value: Essential data.
Decision: Retained.

intensity_counts_raw
Type / values: float column
Written when: Every pixel row
Plain-English meaning: Canonical adapter output before GUI background/smoothing.
Use / value: Essential data/provenance.
Decision: Always retained.

intensity_counts_processed
Type / values: float column
Written when: Only when different from raw
Plain-English meaning: Displayed/result counts after GUI processing.
Use / value: High for direct analysis.
Decision: Conditional; prevents duplicate columns.

# Averaged gated-series fields

file_type
Type / values: str: gated_averaged_series
Written when: Always
Meaning: Selects the matrix-series parser.
Use / value: Essential

sequence_id
Type / values: UUID-like str
Written when: Always
Meaning: Joins output to its acquisition sequence.
Use / value: Essential

mode
Type / values: gated mode str
Written when: Always
Meaning: Planner used to create the series.
Use / value: Essential

timestamp_utc
Type / values: ISO-8601 str
Written when: Always
Meaning: Absolute dating of series completion.
Use / value: Essential

integration_ms
Type / values: float
Written when: Always
Meaning: Applied per-frame exposure.
Use / value: Essential

detector_averages
Type / values: int ≥ 1
Written when: Always
Meaning: Detector samples averaged inside each accepted frame.
Use / value: High

repeat_count
Type / values: int ≥ 1
Written when: Always
Meaning: Requested repetitions of the plan.
Use / value: High

pump_duration_ms
Type / values: int ≥ 0
Written when: When applicable
Meaning: Requested ON time before laser OFF.
Use / value: High

inter_frame_gap_ms
Type / values: int ≥ 0
Written when: When applicable
Meaning: Minimum gap between adjacent acquisitions.
Use / value: High

inter_repeat_gap_ms
Type / values: int ≥ 0
Written when: When applicable
Meaning: Minimum pause between repeat blocks.
Use / value: High

timing_guard_method
Type / values: off/flag/discard
Written when: Always
Meaning: Online timing acceptance policy.
Use / value: High

timing_evaluated_count
Type / values: int
Written when: Always
Meaning: Frames examined by timing guard.
Use / value: High

timing_rejected_count
Type / values: int
Written when: Always
Meaning: Frames excluded by timing guard.
Use / value: High

field_value_mT
Type / values: finite float; zero valid
Written when: Only when Record field checked
Meaning: Associated magnetic field.
Use / value: High when applicable

laser_box_id / laser_channel / laser_wavelength_nm
Type / values: str / int / float
Written when: Laser context available
Meaning: Illumination identity.
Use / value: High

trace.index / label / laser_state / requested_delay_ms
Type / values: int / str / str / int
Written when: Per trace
Meaning: Identifies the state/delay represented by a matrix pair.
Use / value: Essential

trace.sample_count
Type / values: int
Written when: Per trace
Meaning: Accepted spectra accumulated into mean/std.
Use / value: Essential

trace.<timing>.<stat>
Type / values: float, ms
Written when: Per trace
Meaning: For request, call start/mid/end, exposure start/mid/end, uncertainty, and error; each has mean, SD, min, max, median, p95, p99.
Use / value: High/Diagnostic

trace.mean_power_W… / std_power_W…
Type / values: variable float columns
Written when: When power measured
Meaning: Per-channel power mean and spread.
Use / value: High when applicable

wavelength_nm
Type / values: float column
Written when: Every pixel
Meaning: Shared calibrated axis.
Use / value: Essential

<label>__mean_counts / __std_counts
Type / values: float column pairs
Written when: Each output trace
Meaning: Incremental spectral mean and sample SD.
Use / value: Essential

# Fields consolidated, changed, or removed

spectrometer_* rows
Schema-v3 representation: spectrometer key=value row
Reason: Identity values are read together and do not need five vertical lines.

integration_ms + requested_integration_ms
Schema-v3 representation: acquisition.integration_ms + conditional integration_requested_ms
Reason: Applied value is primary; request appears only on a meaningful mismatch.

acquisition_started_s + acquisition_finished_s
Schema-v3 representation: acquisition.call_duration_ms
Reason: Raw perf-counter values are process-local and cannot date a file; elapsed duration is portable.

averaging_mode + device_averaging_used
Schema-v3 representation: acquisition.averaging, or requested/applied pair
Reason: The option itself conveys enabled/applied state; a mismatch remains auditable.

correct_dark + correct_nonlinearity
Schema-v3 representation: detector_corrections row
Reason: Related detector corrections share one line but remain independently explicit booleans.

p_before/p_after/status/command rows
Schema-v3 representation: power row + abnormal-only status row
Reason: Preserves bracketing readings while normal status clutter disappears.

scan_active
Schema-v3 representation: presence of scan group
Reason: Separate feature flag duplicated the existence of scan fields.

background_subtracted
Schema-v3 representation: presence of background group
Reason: Feature is true exactly when a compatible background source is recorded.

background integration/averages/corrections
Schema-v3 representation: removed from output
Reason: Strict compatibility guard requires the background signature to match current applied settings; duplicate values added clutter.

boxcar_width plus implicit enable
Schema-v3 representation: smoothing.method + width + boundary only when active
Reason: Method indicates enabled state and supports future kernels.

gated active flag + separate gated fields
Schema-v3 representation: gated_frame / gated_plan / gated timing groups
Reason: Group presence indicates applicability; related values remain machine-readable.

snr_valid + snr_message
Schema-v3 representation: snr.status + conditional reason
Reason: One status carries valid/invalid state; reason is relevant only on failure.

notes_json
Schema-v3 representation: notes literal CSV field
Reason: JSON encoding duplicated escaping and obscured normal text.

signal_max_counts
Schema-v3 representation: removed
Reason: Exactly derivable from the raw/processed array; storing it risks inconsistency.

laser_port
Schema-v3 representation: removed from spectrum provenance
Reason: COM assignment is transient machine routing, not instrument/sample identity.

raw + processed always
Schema-v3 representation: raw always; processed only if changed
Reason: Eliminates duplicate data columns while preserving both when processing matters.

# Potential combinations deliberately not made

Power before + after → mean only
Decision: Not combined
Why: A mean hides drift and cannot reveal a transient/range change; both readings share one row already.

Detector timing inside acquisition row
Decision: Not combined
Why: Detector/SDK timing is backend-specific and semantically different from controller call duration.

Gated call + exposure timing
Decision: Not combined
Why: Combining them would invite the false assumption that a vendor call equals physical exposure.

Laser + power_scan
Decision: Not combined
Why: Laser identity/setpoint is instrument provenance; requested/expected sample power and filters are scan intent/calibration.

All gated fields on one line
Decision: Not combined
Why: Frame identity, requested plan, observed timing, and quality decision serve separate analysis stages; one line would be unreviewable.

SNR status + all metrics
Decision: Not combined
Why: Status is a quick validity gate; dense metrics are only useful when valid.

Field inside acquisition
Decision: Not combined
Why: A standalone conditional scalar is easy for external scripts to locate and its applicability semantics are unusually important.

Averaged-series timing percentiles
Decision: Not compressed
Why: Wide machine-analysis columns are intentional; eliminating percentiles would discard jitter/tail behavior.

# Metadata that may be valuable but was not added

| APPROVAL | The following are recommendations only. They are not written by this revision, in accordance with the request to seek approval before expanding spectrum metadata. |
| --- | --- |

Exact Andor state snapshot
What it would record: Grating, center wavelength, filter, flipper/ports, focus, offsets, A/D channel, amplifier, H/V speeds, gain, binning.
Why useful: Very high for reproducibility and background compatibility.
Recommendation: Recommend next after confirming names/units; group as andor_state.

Slit width(s) and entrance/exit optics
What it would record: Actual slit state if motorized/queryable.
Why useful: Very high for spectral resolution/throughput.
Recommendation: Add only from readback, never assumed GUI state.

Detector temperature and cooler state
What it would record: Temperature, target, status at acquisition.
Why useful: High for dark current and sensor stability.
Recommendation: Recommend for Andor/QEPro when a fresh or recent read exists; avoid slowing every frame.

Driver/SDK/DLL and software revision
What it would record: App version/commit, SeaBreeze version, SDK2 DLL filename/version.
Why useful: High for troubleshooting and audit.
Recommendation: Recommend compact provenance group; avoid absolute paths/user names.

Background signature/hash
What it would record: Stable detector/grid/settings signature.
Why useful: High for audit after strict guard.
Recommendation: Could expose the existing internal signature without duplicating all fields.

Power-meter identity and channel roles
What it would record: Model/serial, channel mapping, units/range mode.
Why useful: High for multi-channel experiments.
Recommendation: Defer until channel-selection UX is defined.

Magnetic-field source/readback
What it would record: Manual vs plugin/instrument, sensor ID, uncertainty, timestamp.
Why useful: High once fields are measured rather than manually entered.
Recommendation: Add with future field instrument plugin.

Laser actual emission/interlock readback
What it would record: Verified emission state and query age.
Why useful: Safety/diagnostic value.
Recommendation: Only add after bench proves query semantics; do not equate with command latch.

Hardware trigger/exposure timestamps
What it would record: Trigger source, camera FIRE edge/time, clock domain/uncertainty.
Why useful: Essential for higher-accuracy time-resolved work.
Recommendation: Add with hardware-trigger implementation.

Radiometric/instrument-response calibration
What it would record: Calibration ID/version, collection area, correction state.
Why useful: Essential for absolute spectra.
Recommendation: Needs a formal calibration registry first.

Sample/environment/operator metadata
What it would record: Sample ID, temperature, pressure, operator, experiment ID.
Why useful: Potentially high, lab-dependent.
Recommendation: Better supplied by run/session metadata or procedure parameters than repeated ad hoc fields.

Smoothing-kernel parameters
What it would record: Kernel, width/sigma/order/polyorder/boundary mode.
Why useful: High once new smoothing methods ship.
Recommendation: Extend smoothing group alongside the feature, not before.

Plugin/procedure provenance
What it would record: Plugin IDs/versions and procedure definition/hash.
Why useful: High for extensible automation.
Recommendation: Add only with versioned registry/runner.

File integrity checksum
What it would record: Hash of numeric payload or sidecar manifest.
Why useful: Diagnostic/archive integrity.
Recommendation: Consider for archival workflows; self-hash inside the same file is awkward.

# Compatibility and analysis guidance

Treat absent optional groups as “not recorded/not applicable,” not as zero or false.

Use acquisition.integration_ms for normalization. Preserve integration_requested_ms only as diagnostic intent.

Use intensity_counts_raw for reprocessing and intensity_counts_processed for reproducing the GUI result when that column exists.

For gated analysis, filter on gated_quality.status and inspect the timing basis/uncertainty before interpreting a requested-delay axis as actual time. For Andor output, discarded_before means stale at request entry, while discarded_for_target means consumed afterward because its estimated exposure midpoint was too early.

For averaged series, sample_count is the accepted count after timing rejection; it can be lower than repeat_count.

Third-party parsers should ignore unknown key=value names so future schema additions remain forward-tolerant.

# Schema v4 additions to individual spectra

The preceding catalog remains the complete schema-v4 base. The fields below are the approved additions made after the 09/09 catalog was first produced. They are conditional and do not restore redundant feature flags.

software_provenance.gui_version
Type or values: str
Written when: When record is formed
Plain English meaning: GUI semantic version.
Use or value: High: software provenance.
Decision: New grouped field.

software_provenance.source_revision
Type or values: str/date revision
Written when: When record is formed
Plain English meaning: Reviewed source revision.
Use or value: High: source traceability.
Decision: New grouped field.

software_provenance.python
Type or values: version str
Written when: When record is formed
Plain English meaning: Interpreter version.
Use or value: Diagnostic/reproduction.
Decision: New grouped field.

software_provenance.adapter
Type or values: str
Written when: When adapter reports it
Plain English meaning: Adapter implementation ID.
Use or value: High: backend provenance.
Decision: New grouped field.

software_provenance.adapter_version
Type or values: str
Written when: When adapter reports it
Plain English meaning: Adapter revision.
Use or value: High: behavior provenance.
Decision: New grouped field.

software_provenance.vendor_runtime
Type or values: str without absolute path
Written when: When adapter reports it
Plain English meaning: SDK2 DLL filename or SeaBreeze version.
Use or value: High: driver diagnosis.
Decision: New grouped field.

andor_state.grating
Type or values: int
Written when: Andor only, readback available
Plain English meaning: Current Kymera grating.
Use or value: High: wavelength/resolution provenance.
Decision: Approved and added.

andor_state.center_wavelength_nm
Type or values: finite float
Written when: Andor only
Plain English meaning: Kymera center wavelength readback.
Use or value: High: spectral configuration.
Decision: Approved and added.

andor_state.filter_position
Type or values: int
Written when: Andor and wheel present
Plain English meaning: Internal Kymera filter position.
Use or value: High when installed.
Decision: Approved and conditional.

andor_state.focus_mirror_position
Type or values: int
Written when: Andor and focus present
Plain English meaning: Focus mirror step position.
Use or value: High: optical reproduction.
Decision: Approved and conditional.

andor_state.grating_offset
Type or values: int
Written when: Andor and readback available
Plain English meaning: Offset for active grating.
Use or value: High: calibration audit.
Decision: Approved and added.

andor_state.detector_offsets
Type or values: JSON object
Written when: Andor and readback available
Plain English meaning: Detector offsets by port pair.
Use or value: High: calibration audit.
Decision: Approved and conditional.

andor_state.output_amplifier
Type or values: int
Written when: Andor only
Plain English meaning: Selected output amplifier.
Use or value: High: detector response.
Decision: Approved and added.

andor_state.horizontal_readout_mhz
Type or values: finite float
Written when: Andor only
Plain English meaning: SDK2 horizontal clock in MHz.
Use or value: High: noise/throughput.
Decision: Approved; 0.1 MHz is 100 kHz.

andor_state.vertical_shift_us
Type or values: finite float
Written when: Andor only
Plain English meaning: Vertical shift time per pixel.
Use or value: High: CTE/throughput.
Decision: Approved and added.

andor_state.preamp_gain
Type or values: finite float
Written when: Andor only
Plain English meaning: Selected preamplifier gain.
Use or value: High: response/noise.
Decision: Approved and added.

andor_state.horizontal_binning
Type or values: int >= 1
Written when: Andor only
Plain English meaning: Horizontal FVB binning/calibration factor.
Use or value: High: grid/calibration.
Decision: Approved and added.

andor_state.flipper_positions
Type or values: JSON object
Written when: Andor and flippers present
Plain English meaning: Entrance/exit port mirror positions.
Use or value: Analysis/optical routing.
Decision: Optional as answered.

detector_timing_ms.mode
Type or values: single_scan or run_till_abort
Written when: Timing available
Plain English meaning: Native acquisition mode.
Use or value: High: throughput/timing semantics.
Decision: New stream field.

detector_timing_ms.driver_frame
Type or values: int >= 0
Written when: Continuous index available
Plain English meaning: SDK2 frame index returned.
Use or value: High: order/freshness audit.
Decision: New stream field.

detector_timing_ms.discarded_before
Type or values: int > 0
Written when: Continuous backlog removed
Plain English meaning: Frames complete before request and discarded.
Use or value: High for gated freshness.
Decision: Conditional; zero omitted.

detector_timing_ms.discarded_for_target
Type or values: int > 0
Written when: Andor gated stream rejects an early candidate
Plain English meaning: Frames consumed after request because their estimated exposure midpoint preceded the requested gated target.
Use or value: High: distinguishes target alignment from pre-request freshness.
Decision: Conditional; zero omitted and not combined with discarded_before.

detector_timing_ms.pending_after_wait
Type or values: int > 1
Written when: Continuous backlog observed
Plain English meaning: Maximum unread depth after wait.
Use or value: Diagnostic: scheduling/overflow.
Decision: Conditional; ordinary depth omitted.

power.channels
Type or values: JSON array: 1, 2, or both
Written when: Spectrum power measured
Plain English meaning: Physical Newport channels associated with arrays.
Use or value: Essential for channel meaning.
Decision: New; only selected channels saved.

background.signature_sha256
Type or values: hex str
Written when: Background subtracted
Plain English meaning: Strict detector/grid/settings compatibility signature.
Use or value: High: exact match audit.
Decision: Approved and added.

smoothing.boundary
Type or values: nearest, reflect, wrap, constant
Written when: Smoothing active
Plain English meaning: How edges are extended for convolution.
Use or value: High: reproducible processing.
Decision: New; nearest default.

# Current schema v4 averaged gated series

A one-file averaged result now carries the same detector, software, background, processing, field, and selected-power provenance as its accepted source frames. Per-trace power columns name their physical channel explicitly.

file_type / schema_version
Type or values: gated_averaged_series / 4
Written when: Always
Meaning: Payload family and layout version.
Use: Essential

sequence_id / timestamp_utc
Type or values: str / ISO-8601
Written when: Always
Meaning: Sequence join ID and absolute date.
Use: Essential

spectrometer.*
Type or values: backend, model, detector serial, spectrograph serial
Written when: When known
Meaning: Instrument identity.
Use: High

software_provenance.*
Type or values: version strings
Written when: When known
Meaning: GUI, source, Python, adapter, runtime versions.
Use: High

andor_state.*
Type or values: named values
Written when: Andor only
Meaning: Approved readback-confirmed optical/detector state.
Use: High

acquisition.*
Type or values: applied/requested ms, averages, averaging
Written when: Always; request only on mismatch
Meaning: Per-frame acquisition settings.
Use: Essential

detector_corrections.*
Type or values: two booleans
Written when: Always
Meaning: Electric-dark and nonlinearity settings.
Use: High

smoothing.*
Type or values: method, width, boundary
Written when: Only when active
Meaning: GUI processing applied to traces.
Use: High

gated_plan_ms.*
Type or values: mode, repeats, pump, frame gap, repeat gap
Written when: Always/applicable
Meaning: Requested sequence structure.
Use: Essential

timing_guard.*
Type or values: method, evaluated, rejected
Written when: Always
Meaning: Quality policy and result counts.
Use: High

detector_stream.discarded_for_target
Type or values: int > 0
Written when: One-file gated series has target-alignment discards
Plain English meaning: Total early Andor stream frames consumed while selecting all accepted spectra in the series.
Use or value: High: sequence-level timing and throughput audit.
Decision: Conditional detector_stream group; zero omitted.

incomplete_reason
Type or values: str
Written when: Non-user failure with accepted partial data
Meaning: Why output ended early.
Use: High

field_value_mT
Type or values: finite float; zero valid
Written when: Record field enabled
Meaning: Manual field associated with series.
Use: High when applicable

laser.*
Type or values: box ID, channel, wavelength
Written when: When known
Meaning: Excitation identity.
Use: High

power_channels
Type or values: 1, 2, or both
Written when: Power samples retained
Meaning: Physical channels used by trace power columns.
Use: Essential for power

background.*
Type or values: signature, source timestamp
Written when: Background subtracted
Meaning: Exact compatible background identity.
Use: High

trace_metadata_columns
Type or values: column-name list
Written when: Always
Meaning: Schema for each following trace row.
Use: Essential

trace
Type or values: label/state/delay/count, timing statistics, named power columns
Written when: One per state/delay
Meaning: Accepted sample count and distribution summaries.
Use: Essential

wavelength_nm
Type or values: float column
Written when: Every pixel row
Meaning: Shared calibrated wavelength axis.
Use: Essential data

<trace>__mean_counts
Type or values: float column
Written when: Every trace and pixel
Meaning: Incremental accepted-frame mean.
Use: Essential data

<trace>__std_counts
Type or values: float column
Written when: Every trace and pixel
Meaning: Sample standard deviation; zero for one sample.
Use: High

# Schema v4 consolidation changes

All Newport values with ambiguous array position
Current representation: power.channels plus only selected before and after values
Reason: Preserves intended association and removes unrequested channel clutter.

No stable software/adapter record
Current representation: software_provenance named group
Reason: Makes driver and source differences diagnosable without workstation paths.

No persisted Andor optical/readout state
Current representation: conditional andor_state group
Reason: Reproduces detector and Kymera configuration; A/D channel remains omitted by decision.

Single-scan detector timing only
Current representation: detector timing plus mode, driver frame, pre-request discard, target-alignment discard, and pending depth
Reason: Makes circular-buffer order, freshness, and gated frame selection auditable without conflating two different discard causes.

Background timestamp only
Current representation: timestamp plus strict signature
Reason: Proves configuration compatibility without repeating every background setting.

Generic mean_power_W and std_power_W tails
Current representation: mean_power_chN_W and std_power_chN_W
Reason: Names the physical channel and makes the wide trace row self-describing.

Flat averaged-series rows
Current representation: spectrometer, acquisition, gated plan, timing guard, background, and laser groups
Reason: Uses the same compact pattern as individual spectra.

# Fields still not added

Detector temperature, cooler target, cooler state, and reading age remain outside per-frame metadata until query overhead and a maximum acceptable age are measured.

A/D channel remains omitted from Andor spectrum provenance by user decision.

Manual field has no source, uncertainty, sensor identity, or reading age. Add those only with a measured-field plugin so manual and measured values cannot be confused.

Physical OBIS emission/interlock state remains unrecorded because the current query appears to expose command latch rather than proven key state.

Hardware trigger/FIRE edge, clock domain, and uncertainty require a real trigger implementation.

Sample/operator/environment fields remain optional run/procedure concerns; they are not forced into every spectrum.

Close combinations deliberately retained: before and after power are not reduced to a mean; call and exposure timing remain separate; gated identity, plan, observed timing, and quality remain separate groups; raw and processed arrays remain separate only when processing changes the data.