# Spectrum schema v5 supplement

Additions to the retained September 9 metadata reference

Revision: 2026-10-06

| Field / file | Meaning |
| --- | --- |
| intensity_counts_raw | Canonical adapter result; unchanged. |
| intensity_counts_processed | Optional final array after background subtraction and smoothing; unchanged. |
| intensity_counts_unsmoothed | Optional post-background, pre-smoothing values when smoothing is active. Same length/grid as raw data. Written independently of whether raw and final happen to be equal. |
| parameter_scan_json | Conditional JSON dictionary: scan_id, kind, index/count, target/units, elapsed request/completion and settle. Time scans add scheduled_elapsed_s, request_lateness_s, overrun_policy. Temperature scans add temperature_trigger, trigger_age_at_request_s, tolerance/dwell. A drained canceled point adds termination_requested. |
| temperature_trigger | source, value, units, sampled_s, timestamp_utc, status and identity. sampled_s is process-local monotonic evidence; age is independently stored. It is not a hardware-synchronized exposure temperature. |
| *_scan_<uuid>.json | Terminal scan summary: full plan, UUID, UTC start/end, completed/requested acquisition counts, termination, original center and elapsed run time. Written through the ordered executor; not a per-file write receipt or crash journal. |
| Background CSV v1 | Separate file_type=background. Metadata carries signature/configuration and rows contain wavelength_nm, background_counts and background_counts_per_s. Ordinary spectrum loading rejects this file type. |
| Monitor CSV | power_channel and timer_origin_monotonic_s specify the display selection and elapsed origin; both captured physical channel columns are retained. |
| Temperature CSV | Timestamp, elapsed time, source, temperature, units, status and identity for the retained sensor history. |

Compatibility: the new spectrum loader accepts plain legacy spectra and schemas v1–v5. Existing v4 readers that ignore unknown metadata/columns retain the raw/final data, but cannot be assumed to expose v5 additions. External analysis loaders should select columns by header name. Gated averaged-series schema remains v5; file_type separates it from spectrum v5.