# Gated timing and low-light phosphorescence

Field interpretation, fixed-window retention, and sensitivity strategy

Revision date: 2026-09-14

Bottom line The recorded timing families are distinct, but none is a direct optical laser timestamp. The remaining negative values are valid exposure starts for zero-delay windows that straddle the software OFF origin. The new bounded retention mode can collect exactly N spectra that satisfy a recorded timing window, but low-emissivity success will depend more strongly on photon collection, detector settings, and matched background/reference acquisition.

# Direct answers

| Term | Definition |
| --- | --- |
| Request | Coordinator time immediately before emitting the spectrum request into the runtime/worker path. It is not a time inside the spectrometer controller after receipt. |
| Acquisition call | Worker timestamps around adapter acquire_spectrum. Start/midpoint/end bound native wait, transfer, and adapter averaging; they exclude later plotting and saving. |
| Exposure | Best backend estimate of detector integration. Andor dates the stream frame from SDK-applied exposure/readout/cycle information; fallback backends may use call bounds. |
| Timing error | Signed residual: selected exposure reference minus requested delay. Positive = later than requested; negative = earlier. It is not uncertainty. |
| Exposure uncertainty | Conservative half-width/bound on the exposure estimate. It is not standard deviation and excludes unknown command-to-optical laser delay. |
| Laser reference | No optical or hardware timestamp is currently returned by the boxes. The origin follows successful serial state command/readback completion and queued GUI delivery. |

# Are negative values acceptable?

Yes, but only in specific places. At requested delay 0 ms, a finite exposure whose midpoint is after the software OFF origin can begin before it. That exposure-start minimum/mean is physically and mathematically meaningful and must not be clamped. In the supplied eight GUI CSVs, negative means/minima occur only in this zero-delay exposure-start category.

| Observation | Interpretation | Action |
| --- | --- | --- |
| Exposure start < 0 at 0 ms midpoint target | Usually acceptable | Exposure straddles the transition origin; check midpoint/end and basis. |
| Timing error < 0 in fixed midpoint/start policy | Rejected by strict window if outside −tolerance | Selected reference was earlier than requested. |
| Request or call start < 0 | Investigate | Request/work apparently preceded the stored origin; may indicate ordering/reference error. |
| Exposure midpoint < 0 when policy requires ≥0 | Investigate/reject | Selected frame does not satisfy the target policy. |
| Any value with controller-call-bounds basis | Lower-confidence estimate | It is a software bound, not detector telemetry. |

# What the 2026-09-14 evidence shows

| Run | Frames | Backlog | Target | Stale | Call median | Midpoint residual |
| --- | --- | --- | --- | --- | --- | --- |
| Unaligned stream-gap | 100/100 | 595 | 0 | 99 | 4.367 ms | −16.562 ms; all 99 early |
| Midpoint-aligned stream-gap | 100/100 | 594 | 198 | 99 | 34.353 ms | +13.425 ms; all 99 eligible |

The first request takes about 1.05–1.10 s because it prepares/arms SDK2; GUI gated acquisition consumes an unsaved preparation frame before laser timing.

Aligned calls take longer because the selector discards early frames and waits for the first eligible midpoint; it does not re-arm the camera.

Fixed 100-frame GUI traces show midpoint timing-error means near 6–7 ms. One-sample interleaved traces spanned about 0.14–31.08 ms; ten-sample traces about 3.77–11.39 ms.

The background-subtracted Rhodamine B data remain close to zero in the tested geometry. That is evidence of inadequate detected signal under those conditions, not evidence that timing selection alone can recover the phosphorescence.

# Retain N inside fixed window

1.  Choose delayed-after-OFF or interleaved-decay and One averaged series file.

2.  Set Repeats to the number of accepted spectra required for every trace.

3.  Select exposure midpoint (recommended first) or exposure start, tolerance, and attempt multiplier.

4.  The lazy preflight computes the accepted target and finite maximum attempts, warns when the window is narrower than detector cadence, and rejects incompatible Newport/duplicate-delay settings.

5.  For each pump/OFF cycle, the Andor selector waits from the lower edge of the requested window; the coordinator makes the authoritative two-sided decision.

6.  Rejected arrays are released. Accepted arrays update incremental mean and sample variance. Completed traces are skipped in later attempts.

7.  The run stops at exactly N accepted spectra per trace. If the ceiling is reached, accepted partial traces are saved with an incomplete reason and full attempt counts.

| Setting | Meaning |
| --- | --- |
| Requested delay | 10 ms |
| Reference | Exposure midpoint |
| Tolerance | ±2 ms → accepted interval 8–12 ms |
| Repeats | 100 accepted spectra required |
| Attempt multiplier | 3× → maximum 300 detector attempts for one trace |
| Stop condition | 100 accepted, or finite ceiling failure with partial output |

Important: a ±2 ms software window is narrower than the approximately 15 ms free-running iDus period measured here. The mode is feasible, but yield can be low and the retained software window does not prove an optical ±2 ms window. Start wider (for example ±8 ms), measure yield and optical offset/jitter, then narrow only if supported.

# Sensitivity options ranked by likely impact

| Rank | Technique | Implementation | Potential | Constraint |
| --- | --- | --- | --- | --- |
| 1 | Optical collection / étendue | Re-image the emitting volume onto the entrance slit; maximize accepted NA; check cryostat-window clipping; widen slit while resolution allows; remove unnecessary filters/fiber losses. | High | Usually the largest gain; must measure spectral-resolution tradeoff. |
| 2 | Spectrograph configuration | Choose grating blaze/center and output port for emission band; verify focus/calibration; use the lowest-loss path and appropriate order sorting only when necessary. | High | Immediate, reversible after focus contract is resolved. |
| 3 | Pump/collection duty cycle | Lengthen pump to populate triplets, acquire as early as scatter permits, collect multiple useful OFF windows per pump, and minimize dead time. | High | Sample photostability/heating can limit. |
| 4 | Detector operating point | Cool and stabilize; compare supported preamp gain and horizontal/vertical rates using blocked-light plus weak-source measurements; lengthen exposure and average accepted cycles. | Medium–high | Read-noise benefit depends on photon/background regime. |
| 5 | Matched backgrounds | Use the same applied exposure/readout/grating/center/port for background; pair nearby reference cycles to reject drift; use robust cosmic-ray rejection. | Medium–high | Essential when signal is a small difference. |
| 6 | Phase-cycled boxcar | Alternate signal/reference phases, subtract paired spectra, aggregate mean and propagated variance synchronously with pump state. | Medium | Rejects DC/drift; reference adds noise and does not create photons. |
| 7 | Band/template estimator | Integrate a predeclared wavelength band or project onto a validated spectral template with inverse-variance weights; preserve raw spectra. | Medium | Can improve detection over peak pixels; avoid post-selection bias. |
| 8 | Post-hoc denoising | Use carefully validated smoothing/baseline models for visualization/secondary analysis, never as a substitute for raw evidence. | Low–medium | Can distort weak broad signals if overfit. |

# Software lock-in: feasible form

A CCD cannot reproduce a conventional analog lock-in merely through software. The feasible analogue is synchronous phase cycling or boxcar detection. For each cycle, acquire a signal phase and a matched reference phase, compute their difference, and average the paired differences. Static background and slow drift cancel to first order; independent noise from both phases remains.

Phase definitions: pump-then-delay signal versus pump-blocked/no-pump reference, or early versus late OFF windows when the late window contains negligible emission.

Data product: per-pixel paired difference, running mean, sample variance, phase timestamps/settings, rejected-pair reason, and optional band/template scalar with uncertainty.

Scheduling: host coordinator owns alternating phases, laser safety, detector stream/session, cancellation, and bounded retries. Settings and background objects are prepared before timed actions.

Limitation: if both frames are independent and similarly noisy, subtraction increases variance; the method wins only when rejected background/drift exceeds that penalty.

Upgrade path: a hardware modulator/trigger and detector synchronization would provide a more conventional lock-in reference and tighter timing than serial software control.

# Recommended first measurement protocol

1.  Use a bright stable lamp or fluorescent sample to optimize cryostat-to-slit image, slit width, grating, center wavelength, and port while recording throughput versus resolution.

2.  With the optical path blocked, measure detector mean/variance/cosmic-ray rate across supported gain and speed combinations at the intended exposure and temperature.

3.  Measure an excitation-on blank/matrix and a pump-off background with identical settings to identify scatter, fluorescence, and drift.

4.  Use Rhodamine B only as an alignment/visible-emission control; verify a clear detected band before attempting weak phosphorescence.

5.  Run fixed-delay retention at 10 ms, midpoint reference, 20 accepted, ±8 ms, 3×. Inspect yield and timing distributions; then narrow if justified.

6.  Add a photodiode timing measurement before interpreting absolute delay. Record command-to-optical offset/jitter for the selected laser channel.

7.  Only after a reproducible nonzero signal, increase accepted repeats and compare ordinary averaging against paired phase-cycled subtraction using a predeclared band/template metric.

# References and evidence boundary

Andor SDK2 Programming Guide: Run Till Abort, GetNumberNewImages, GetOldestImage, acquisition timing, and native return-code behavior. https://andor.oxinst.com/downloads/uploads/Andor_Software_Development_Kit_2.pdf

Project evidence: andor_probe_results_20260914.zip (13 JSON probes, eight complete gated CSVs), plus the cumulative September 9–11 review artifacts.

All optical-collection and synchronous-detection recommendations are general experimental-design guidance. Their quantitative benefit must be measured on the installed cryostat/sample/laser/detector geometry.