# Next-step questions

Focused decisions for Andor performance, metadata, cooler safety, and extension design

| STATUS Reviewed source design | REVISION 2026-09-10 | AUDIENCE Lab users and maintainers |
| --- | --- | --- |

| USE | Type beneath each prompt or highlight an option. Recommended defaults are marked; questions are ordered by the next implementation dependency. |
| --- | --- |

# A. Andor timing evidence

## A1 — Timing probe result

Close Solis, block detector light, leave cooler state unchanged, run the command, and attach the JSON.

python -m troubleshooting.andor_timing_probe --camera-dll atmcd64d_legacy.dll --vertical-speed-index 0 --vertical-speed-index 1 --exposure-ms 15 --frames 50 --output andor_timing.json

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## A2 — Exposure versus frame period

If SDK2 reports 5 ms applied exposure plus a 13.26 ms minimum frame period, should the GUI allow and save 5 ms while showing the throughput limit separately?

Options / recommendation: Recommended: Yes—these are different physical quantities. Alternative: constrain exposure to frame period for simpler but less precise semantics.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## A3 — Readout profile

Should the GUI default to ~16.25 µs/pixel or expose named profiles?

Options / recommendation: Recommended: Quality (16.25 µs) default plus Fast (8.25 µs) after both are characterized. Other: always quality / always fastest.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## A4 — Continuous acquisition scope

If Run-Till-Abort is substantially faster, where should it be enabled first?

Options / recommendation: Recommended: Live first, then gated after trigger/buffer timing validation. Other: Live and ordinary rapid acquisition / gated immediately.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

# B. Cooler shutdown safety

## B1 — Closing during warm-up

Should normal close be blocked during guided warm-up?

Options / recommendation: Recommended: Block, show temperature/progress, and provide an explicit confirmed emergency exit that states warm-up was incomplete.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## B2 — Safe threshold

Is ≥0 °C the normal shutdown threshold for both iDus cameras?

Options / recommendation: List any camera-specific threshold, ramp, or required hold time.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## B3 — Recovery marker

Should startup warn if the prior session ended with cooler ON or temperature below threshold?

Options / recommendation: Recommended: Yes; store only a local incomplete-shutdown marker and query actual state on reconnect.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

# C. Metadata approval

## C1 — Andor state snapshot

Approve queried grating, center wavelength, filter, flipper/ports, focus, offsets, A/D channel, amplifier, H/V speeds, gain, and binning per spectrum?

Options / recommendation: Recommended: Yes, as one conditional andor_state group with readback-confirmed values only.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## C2 — Detector temperature

Approve recent temperature, cooler state, target, and reading age without a new high-rate query per frame?

Options / recommendation: Recommended: Yes; update on a low-rate timer only while relevant.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## C3 — Software/driver provenance

Approve GUI revision, adapter version, and vendor DLL/SeaBreeze version without absolute paths?

Options / recommendation: Recommended: Yes, in one compact provenance group.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## C4 — Background signature

Expose the internal strict compatibility signature/hash?

Options / recommendation: Recommended: Yes; this audits matching without duplicating background settings.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## C5 — Run metadata

Should sample ID, operator, sample temperature, and experiment ID be entered once per run/session?

Options / recommendation: Recommended: Add a small optional run form only if these are routinely used; otherwise keep notes.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

# D. Gated quality policy

## D1 — Outlier disposition

After characterization, should the normal default remain Discard or change to Flag?

Options / recommendation: Recommended: Discard for averaged output; Flag for individual files where preserving evidence matters.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## D2 — Baseline grouping

Should all laser-OFF frames share one baseline or should each delay/state have one?

Options / recommendation: Current: one OFF baseline. Recommended if bias depends on delay: separate by adapter and gated label/delay family.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## D3 — Stable absolute bias

Should preflight enforce an absolute maximum delay bias as well as outlier detection?

Options / recommendation: Recommended: Warn but do not discard until hardware timing defines an acceptable accuracy bound.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## D4 — Aborted output

Should accepted partial averaged data be saved after timing-abort?

Options / recommendation: Recommended: Save an explicit _ABORTED file with reason/counts; never name it as complete.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

# E. Power and laser behavior

## E1 — Newport channels

Should spectrum-associated power support channel 1, channel 2, or both independently?

Options / recommendation: Recommended: Multiselect with channel roles/names; require one only when Measure power is checked.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## E2 — Wrong-channel guard

When a selected channel is dark while another sees signal, should the GUI warn or block?

Options / recommendation: Recommended: Warn after a configurable consecutive count; do not block automatically.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## E3 — OBIS physical key

Can you test whether emission-state readback changes when the physical key/interlock opens while command state remains ON?

Options / recommendation: The result determines whether conditional polling can detect actual desynchronization.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

# F. Plugins and procedures

## F1 — Trust model

Who can install and enable plugins?

Options / recommendation: Recommended: Maintainer installs; users enable from an approved list; visibly mark unsigned development plugins.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## F2 — First pilot

Which capability should prove the API?

Options / recommendation: Recommended: Read-only magnetic-field sensor—small surface, immediate metadata value, no laser actuation.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## F3 — Procedure v1

Should v1 be declarative YAML/JSON with bounded loops/conditions and no arbitrary Python?

Options / recommendation: Recommended: Yes.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## F4 — Authoring interface

Which first authoring UX is preferred?

Options / recommendation: Recommended: Schema-aware text editor plus parameter form and emulator dry-run preview; visual builder later.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## F5 — External API

Is normal VS Code/Jupyter automation more valuable than an integrated Python IDE?

Options / recommendation: Recommended: Local API first; avoid an unsafe embedded runtime.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## F6 — Remote instruments

Must extensions support instruments on another computer within the next year?

Options / recommendation: This determines local-only versus network-ready first RPC design.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

# G. Compact acquisition layout

## G1 — Advanced section

Confirm always-visible Exposure, Averages, Smoothing, Field, Live, and Take Spectrum; move corrections, averaging backend, background, and power association to Advanced.

Options / recommendation: Recommended from the answered document; list anything that must stay visible.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## G2 — Smoothing preview

Should the tiny inline kernel preview open a larger popup on click?

Options / recommendation: Recommended: Yes; inline has no axes, popup shows normalized weights and parameters.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

## G3 — Initial kernels

Confirm None, Boxcar, Hann, Hamming, Gaussian; Savitzky–Golay later.

Options / recommendation: Also specify preferred boundary behavior (reflect/nearest/trim) and Gaussian sigma convention if already decided.

Answer: ___________________________________________________________________________

______________________________________________________________________________________________

# Additional context or priorities

______________________________________________________________________________________________

______________________________________________________________________________________________

______________________________________________________________________________________________