# Remaining implementation questions

Bench evidence and extension decisions after the 2026 09 10 timing review

| STATUS Reviewed source design | REVISION 2026-09-10 | AUDIENCE Lab users and maintainers |
| --- | --- | --- |

The current source implements the answered 09/10 choices. These questions cover only evidence or design decisions that still change production behavior. Complete the bench sections first; the extension sections can wait.

# Andor bench evidence

## A1 Short exposure matrix

Run the command below with Solis closed and detector light blocked, then attach the JSON. Did SDK2 apply 1, 5, or 10 ms in Run-Till-Abort, and did any mode return an error?

Options and recommendation  Recommended: keep requests available but claim only the applied values measured here.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

python -m troubleshooting.andor_timing_probe --camera-dll atmcd64d_legacy.dll --vertical-speed-index 0 --vertical-speed-index 1 --exposure-ms 1 --exposure-ms 5 --exposure-ms 10 --frames 100 --output andor_short_exposure.json

## A2 Gated stream semantics

For requested delays shorter than one frame period, is the accepted policy still the first frame that completes after the request even when its exposure may have begun before the requested delay?

Options and recommendation  Recommended: yes for the current software-timed mode, because every file records the exposure estimate and uncertainty. Future hardware trigger mode can require exposure start after the edge.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

## A3 Cooler threshold hold

Should normal close occur immediately on the first cooler-OFF reading at or above 0 °C, or must temperature remain above the threshold for a hold period?

Options and recommendation  Current implementation: first reading at or above 0 °C. Alternative: require three consecutive readings or a 30-second hold if Andor guidance or lab practice calls for it.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

# Gated quality characterization

## B1 Production guard defaults

After several stable and deliberately loaded runs, what rejected fractions and stable residual biases are observed for QEPro and Andor in each gated mode?

Options and recommendation  Recommended: attach representative CSVs before changing 4.5 sigma, five warm-up samples, or abort above 25 percent after ten evaluations.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

## B2 Stable bias warning threshold

Should a future absolute-bias warning use one common millisecond threshold, a detector-specific value, or a fraction of exposure/frame period?

Options and recommendation  Recommended: detector and gated-mode-specific characterization, then warn from a physically meaningful limit without discarding solely for stable bias.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

# Procedure authoring

## C1 Procedure source format

Which source should the first editor save? Both choices compile into the same typed internal plan and use the same generated form and dry-run preview.

Options and recommendation  Recommended: restricted YAML as the human format plus JSON as canonical interchange. Alternatives: JSON only for strict simplicity, or support both equally from v1 at greater test cost.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

## C2 Expression scope

What calculations must v1 procedures perform beyond named parameters, units, bounded loops, simple comparisons, and references to earlier read-only results?

Options and recommendation  Recommended: no general expressions in v1; add a tiny typed expression language only after concrete procedures demonstrate the need.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

# Administrative and API interaction

## D1 Administrative mode

Should plugin installation require the explicit --admin launch flag even when the Windows account is an administrator?

Options and recommendation  Recommended: yes. OS permission controls whether files can be written; --admin is explicit application consent and enables install/remove UI for that session only.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

## D2 First external API model

Which first interaction is more valuable: start the normal GUI and connect a notebook to its localhost API, or run a separate headless measurement service that the GUI and notebooks both use?

Options and recommendation  Recommended: GUI-hosted localhost pilot for read-only and emulated work; design messages so they can move to a headless service for remote instruments later.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

## D3 Remote deployment horizon

Within the next year, do remote instruments need only read-only monitoring, or must remote clients run laser and acquisition procedures?

Options and recommendation  Recommended: read-only remote pilot first. Time-sensitive and hazardous sequences should execute entirely on the instrument host under local policy.

Answer  ________________________________________________________________________

____________________________________________________________________________________________

# Optional metadata

## E1 Detector temperature sampling

If a benchmark shows negligible overhead, what maximum reading age is acceptable when recording Andor temperature on slow spectra?

Options and recommendation  Recommended: query only when exposure or total averaging time exceeds one second, cache the result, and record temperature plus reading age. Do not add it until the benchmark and age are approved.

Answer  ________________________________________________________________________

____________________________________________________________________________________________